# How to Calibrate a Micrometer: Instrument Calibration Log

This calibration log documents how to calibrate a micrometer for accurate readings, using a step-by-step zero-alignment procedure. The record covers a 0–1 inch Fowler micrometer and a 1–2 inch STORM micrometer, and it can be adapted for other mechanical micrometers of similar design. Use this log to capture instrument identity, the standards applied, the results of each calibration check, and the final release decision.

# Instrument Details

| Instrument ID | Model | Serial Number | Location | Owner | Intended Use |
| --- | --- | --- | --- | --- | --- |
| Not specified | Fowler 0–1" 0.0001" micrometer | Not specified | Workbench | Not specified | Precision linear measurement |
| Not specified | STORM 1–2" 0.0001" micrometer | Not specified | Workbench | Not specified | Precision linear measurement |

Lay out the micrometer, its calibration tool (spanner wrench), and any included standards or accessories before beginning work.

Micrometers and calibration tools laid out on a workbench, with cases open and accessories visible

Note: Instrument ID, serial number, and owner fields should be completed from your lab's asset records; the demonstrated procedure does not assign these identifiers.

# Calibration Schedule

| Frequency | Due Date | Procedure | Responsible Role |
| --- | --- | --- | --- |
| Periodic recheck recommended | Not specified | Zero-alignment calibration using ratchet stop and spanner wrench | Not specified |

The procedure recommends rechecking calibration periodically to maintain accurate readings, but it does not specify an exact interval, a due date, or a named responsible role. These details should be defined by your lab's calibration policy and entered here for each recurring log entry.

# Standards Used

| Standard | Description | Lot Number | Certificate | Expiration Date | Traceability |
| --- | --- | --- | --- | --- | --- |
| Spanner wrench | Calibration adjustment tool included in the micrometer kit | Not specified | Not specified | Not specified | Not specified |
| 1" gauge block | Cylindrical reference standard stored in the micrometer case, used to zero micrometers with a range greater than 0–1" | Not specified | Not specified | Not specified | Not specified |

For micrometers larger than 0–1", locate the reference standard included in the kit (for example, a 1" gauge block for a 1–2" micrometer). The reference standard is typically a cylindrical metal piece stored in the micrometer case.

A STORM 1–2 inch micrometer in its case, with the reference standard and spanner wrench visible; a hand is reaching for the standard

Lot numbers, certificates, expiration dates, and formal traceability records are not addressed in the demonstrated procedure and should be added from your standard's certificate of calibration.

# Calibration Results

## Standard-range micrometer (0–1")

Close the micrometer using the ratchet stop. Hold the micrometer and gently close the spindle until the anvil and spindle faces touch, then use the ratchet stop for final tightening to ensure consistent pressure.

A hand holding a Fowler micrometer, using the ratchet stop to close the spindle completely

Check the zero alignment. Inspect the scale to confirm the '0' on the thimble aligns exactly with the '0' line on the sleeve. If the zeros do not align, note the offset (for example, two thousandths off).

Close-up of the micrometer scale showing the zero marks slightly misaligned

Insert the spanner wrench. Locate the small adjustment hole on the sleeve, usually on the backside of the micrometer, and insert the tip of the spanner wrench so the sleeve can be rotated for calibration.

A hand inserting the spanner wrench into the adjustment hole on the micrometer sleeve

Adjust the sleeve. Gently rotate the sleeve using the spanner wrench until the zero marks align, then remove the tool and recheck alignment by closing the micrometer again. Repeat the adjustment if the marks are still slightly off.

A hand using the spanner wrench to make fine adjustments to the micrometer sleeve

Fine-tune the adjustment. Continue making small adjustments, checking the zero alignment after each change; if you overshoot, adjust in the opposite direction.

Hands making final fine adjustments to the micrometer while checking the zero alignment

Verify the result. Once the zeros are perfectly aligned, the micrometer is calibrated and ready for accurate measurements.

Hands holding the micrometer, confirming the zero alignment after adjustment

| Measured Value | Acceptance Limit | Pass/Fail | Analyst Initials |
| --- | --- | --- | --- |
| Zero offset observed, then adjusted to zero alignment | Zero marks must align exactly | Pass after adjustment | Not specified |

## Larger-range micrometer (1–2", 2–3", etc.)

For micrometers with a range greater than 0–1 inch, use a reference standard included in the kit (such as a 1" gauge block) in addition to the steps above.

Insert the reference standard. Remove the standard from the case and place it between the anvil and spindle; this sets the correct zero point for larger micrometers.

A hand holding the 1 inch reference standard above the open jaws of a STORM 1–2 inch micrometer, preparing to insert it for calibration

Zero the micrometer using the reference standard. Close the micrometer gently on the standard using the ratchet stop for consistent pressure, then check the zero alignment on the scale. If the zeros do not align, use the spanner wrench to adjust the sleeve as described above, repeating until the zero marks align with the standard in place.

Confirm alignment. After adjusting the sleeve, ensure the zero marks on the thimble and sleeve are perfectly aligned so future measurements will be accurate.

Hands holding a micrometer showing the zero marks perfectly aligned after calibration, with a workbench of micrometer cases and tools in the background

| Measured Value | Acceptance Limit | Pass/Fail | Analyst Initials |
| --- | --- | --- | --- |
| Zero offset checked against reference standard, then adjusted to zero alignment | Zero marks must align exactly with the reference standard in place | Pass after adjustment | Not specified |

# Maintenance

Once calibration is complete, return the reference standard and spanner wrench to the case for safekeeping.

Both the STORM 1–2 inch and Fowler 0–1 inch micrometers on the workbench, with their cases open and all calibration tools and standards visible

Preventive maintenance activities, repairs, parts replaced, and any external service provider are not addressed in the demonstrated procedure and should be logged separately when they occur.

# Out-of-Tolerance Actions

If the zero marks do not align after closing the micrometer, note the offset and repeat the adjustment: insert the spanner wrench into the adjustment hole, rotate the sleeve, remove the tool, and recheck alignment, repeating as needed until the zeros align. For larger micrometers, repeat this adjustment with the reference standard in place until the zero marks align with the standard.

Impact assessment on prior measurements, review of affected data, and any quarantine of the instrument are not addressed in the demonstrated procedure; these should be evaluated according to your lab's out-of-tolerance protocol before the instrument is returned to use.

# Approval

Reviewer, approval date, and release-to-use decision are not addressed in the demonstrated procedure. This section should be completed by the designated reviewer once the calibration results above are verified and the instrument is confirmed fit for use.